## Slide 1: Load Pull Application

## Slide 2: What I want the app to do
Calibrate
  Plot calibration as it runs
  Plot verification
  Save calibration data
  Recal cal data for pertinent tests
Collect data
  Plot live data as tests run
  Capability for warning/shut off if device fails to protect device
Plot data after the fact for comparison

## Slide 3: Questions
How to have cal, for power still valid for PA? (need to think about dynamic range)
  How do I determine noise floor of receiver: they say -120 dBm but I think that the measurement environment is noisier than that so I don’t think that is true for this set up
  -10dBm receiver saturation.
  Ask Rhode: how to measure s22 of a PA
Ask Brett for microscope

## Slide 4: Calibrate

## Slide 5: How it works
Steps:
  -calibrate:
Maps to a calibration key
Saves the cal key during the ‘save’ step to a cal file loaded to cal_store
Calibrations are saved to the bench. Multiple cal tests could over write the same variable so make sure to be careful with naming scheme
To reference cal from other files use ${cal.<name>} in the step:
  i.e.

## Slide 6: Power cal
Need:
  Cal1.cal -> outer reference
  Cal2.cal -> probe tip reference
  PM.s1p
How to validate power cal:
  Test once for constants. Rerun to validate:
  Input power = coaxial – probe IL
  Input power = output power + probe cal + thru loss (pretty much 0)
  Input power = output power + tuner connection loss + power meter measurement

## Slide 7: Power Cal Steps
SOLT cal output plane:
  Save file to VNA
SOLT cal probe tip:
  Save file VNA
Save S1p file for PM S11
Save/find s2p file for thru
Code:
  Calibrate
    Load outer plane cal
      Get error terms
    Load inner plane cal
      Get error terms
    Calculate output probe power
    Calculate input probe power
    Save a1^2 coupling to input power
    Save b2^2 coupling to output power

## Slide 8: Notes
Rhode Documentation says changing b2 attenuation invalidated calibration. Looking at validation, load impedance detection changes from -16dBm error to -7 dBm error. It is not my favorite thing obviously but it could be a possible path forward depending on if compression: I think this error is worse.

## Slide 9: Plotter

## Slide 10: Getting Started
cd plotter
pip install -e .
lp-plotter

## Slide 11: What is the data?
Hard coded in some spots right now to be in reflown/runs

## Slide 12: Database

## Slide 13: Sqlite3 database
To open in visual studios:
  Must have sqlite extension
  In searchbar type:
    >sqlite:open database
If things are being weird -> delete database and run populate_db.py
  python -m plotter.populate_db --root plotter --db runs/plotter_database.sqlite
  Must be frun from \reflown not from \reflown\plotter

## Slide 14: Populate_dB
This should be good one programs are set (do not touch test-types after final launch)
Populate loads the columns from the LAST test run under the name

## Slide 15: Source Code

## Slide 16: Flowchart
Start
Bench
Test
Session
Init JsonWriter
Plot?
Cal?
Meta: benchtype, instruments etc
Sequence: Initialize, Load env params load run actions
Action Type?
Error or Ctrl+C?
Fail / Interupt Policy
End
Bench Cal
Load store_cal (persists)
Run shutdown
None
Yes
No
Yes
Yes
No
No
Yes
Initialize plot and plot righter

## Slide 17: YAML

## Slide 18: Different steps:
Calibrate
  Saves variables to persistent (for testbench) variables to be used in code
    i.e. saves input and output coupling for power calibration for VNA
  Set limits through cal.
Measure
  Calls a function/method for a specific instrument that returns a value that can be saved
    i.e. get_voltage for DMM
Call
  Calls a function/method for specific instrument that does not return a value
Sweep
  Will sweep over a variable. Can be nested i.e. power and frequency
  Can subcall call/measure. Unsure about calibrate.
results_update or update_results
  Updates flattened environment to results
  If plot is defined it will update plot
  Will check if limits are reached and if shutdown needs to be run. This is the only way to check if limits have been met
plot_reset
  Saves a png and then reinitializes the plot
Transform
  Runs method/function that can return or not return a value but is not connected to an instrument. Passes in current environment variables and calibration

## Slide 19: Transforms

## Slide 20: Backend transforms: Don’t worry about it tbh
_read_sweep_points
  helper to read mag/deg rows from a CSV into a list of point tuples (used by load-sweep transforms).
_power_correction_cal
  computes coupling factor and related values using DUT/PM calibration terms, PM power, and S-parameter data; returns C10_2 and supporting vectors.
_extract_frequency_vector
  best-effort extraction of a frequency array from dict/list/ndarray structures.
_frequencies_overlap
  checks whether two frequency arrays share the same values within a tolerance.
_extract_scalar
  pulls out a single numeric value from mixed input types (scalars, strings, dicts, objects, arrays).
_convert_dbm_to_linear
  converts a dBm value to linear watts.
_extract_array_field
  extracts a numeric array from dict/object fields matching a set of keys, with optional type conversion.
_to_array
  normalizes input (including real/imag dicts) to a NumPy array of the requested dtype.
_align_array_length
  broadcasts a length-1 array to a target length or returns None if it can’t match the size

## Slide 21: Make new transform
Define in transform.py
Transform.py is a registery. Once function is defined add it to the registry with registry.register
Example:
  Def foo(payload: dict, _cal: dict) -> dict:
    var = payload.get(“var") or {}
    var_return = 2*var
    Return {“var_name”: var_return }
  registry.register(“foo_name", foo)

## Slide 22: VNA attenuation test
Does changing Attenuation maintain cal?

## Slide 23: Generating Load Sweeps
It’s the same for source just a different word

## Slide 24: Sweep geneation
python -m src.loadpull.tools.make_load_sweep
  Generates sweep in csv file with each point at a different impedance. Impedances are single frequency specific
Make_load_sweep parameters:
  --center-mag Center |Gamma| for the rings (0..1).
  --center-deg Center angle in degrees for the rings.
  --radii Comma-separated ring radii relative to center (e.g., `0.0,0.1,0.2`); `0.0` is added if omitted.
  --angle-start / --angle-stop Start/stop angles (deg) for each ring (default -180/180).
  --resolution Angular preset per ring: `coarse` (~18 pts at outer ring), `fine` (~36), `ultra` (~54). Ignored if points-per-ring or deg-step is provided.
  --points-per-ring Fixed number of points for each nonzero ring (overrides resolution).
  --deg-step Fixed angular step in degrees for all rings (overrides resolution if given).
  --out Output CSV path.
Example:
  python -m src.loadpull.tools.make_load_sweep --center-mag 0.25 --center-deg -90 --radii 0.25,0.5 --angle-start -180 --angle-stop 180 --resolution coarse --out loadsweeps/load_sweep_attenvalidation.csv

## Slide 25: Using Sweeps
Sweep file is used as parameter in YAML:
  parameters: {freq_ghz: {default: 5}, sweep_file: {default: "loadsweeps/test_Tuner_sweep.csv"}}
Sweep is loaded into yaml:
  steps:
  - transform:
  method: loadpull_sweep_len
  args: { file: "${sweep_file}" }
  save_as: "sweep_info“
Sweep is swept using index
- sweep:
var: idx
from: 0
to: "${sweep_info.count}"
step: 1
do:
- transform:
method: loadpull_sweep_point
args: { file: "${sweep_file}", index: "${idx}" }
save_as: "point"
- call: { inst: LOADTUNER, method: set_gamma, args: ["${freq_ghz}", "${point.gamma_mag}", "${point.gamma_deg}"] }

## Slide 26: Plotter

## Slide 27: Plotter Registry

## Slide 28: Plotter Registry
The plotter Registry is where the different types of plots for each test type is defined.
Each test type can have multiple tabs
Each plot is defined by:
  title: plot title.
  mode: line or scatter.
  x: x-axis field (flat key).
  y: list of y fields.
  xy: tuple of two fields for scatter (x,y) pairs.
  columns: list of columns required.
  xlabel/ylabel: axis labels.